## Slide 1
Ventilator Control Architecture
Three-Layer Control System for Medical Ventilators
Åke

## Slide 2
System Overview
Layer 3: High-Level Controller (HLC)
8-state ventilation mode control
Layer 2: Low-Level Controller (LLC)
Breathing phase control (insufflation/exsufflation)
Layer 1: PID Controllers
Actuator control (O2, Air, Exp Valve, Blower)
Control Philosophy: Indirect control through setpoint manipulation

## Slide 3
Layer 1: PID Controllers
O2 Flow
Controls oxygen flow rate
Air Flow
Controls air flow rate
Exp Valve
Controls expiratory pressure
Blower
Controls blower flow rate
Characteristics:
Independent operation
Sensor scaling and linearization
Receives setpoints from LLC
Fast, stable control loops

## Slide 4
Layer 2: Low-Level Controller
Insufflation State
Delivers gas to patient
Controls O2 and Air flow
Maintains insp pressure
Exsufflation State
Releases pressure
Controls exp valve
Maintains PEEP
State Transitions:
Patient effort detection (flow/pressure triggers)
Setpoint changes from HLC
Pressure limiting for safety
LLC → HLC Communication:
State changes, measured tidal volume, pressure limit events

## Slide 5
LLC Parameters (Set by HLC)
FiO2
Fraction of inspired oxygen (0.21-1.0)
Insufflation Pressure
Target pressure during inspiration
Pressure Limit
Maximum allowed airway pressure
Exsufflation Pressure (PEEP)
Target pressure during expiration
Insufflation Flow
Target flow rate during inspiration
Bias Flow
Continuous flow during expiration
Max Inspiratory Flow
Maximum allowed inspiratory flow
Flow Trigger Threshold
Flow change to detect patient effort
Pressure Trigger Threshold
Pressure change to detect patient effort
HLC controls LLC behavior by manipulating these setpoints

## Slide 6
HLC Parameters
Timing Parameters:
Respiratory Rate (RR) - breaths per minute
I:E Ratio - inspiration to expiration time ratio
Inspiratory Pause Fraction - plateau time as fraction of cycle
State Durations - time for each of the 8 states (can be zero)
Pressure Parameters:
Inspiratory Pressure - baseline during inspiration
PEEP - baseline during expiration
Insp Support Pressure - boost when patient triggers in insp support
Exp Support Pressure - boost when patient triggers in exp support
Pressure Limit - maximum safety limit
Flow & Volume Parameters:
Tidal Volume (Vt) - target volume per breath
Max Inspiratory Flow - flow rate limit
Bias Flow - continuous background flow

## Slide 7
Layer 3: High-Level Controller
8-State Sequential State Machine
Inspiration Sequence
1. Non-triggerable
Fixed mandatory breath
2. Support
Patient can trigger pressure support
3. Synch
Patient-synchronized transition
4. Pause
Inspiratory plateau
Expiration Sequence
5. Non-triggerable
Fixed expiration phase
6. Support
Patient can trigger support
7. Synch
Patient-synchronized transition
8. Pause
Expiratory pause

## Slide 8
Tidal Volume Control
Initial Flow Calculation
Insufflation Flow = Vt / Active Inspiration Time
Breath-to-Breath Adaptive Control:
1. Measure actual delivered tidal volume (Vt_measured)
2. Calculate error: ΔVt = Vt_target - Vt_measured
3. Gradually adjust flow setpoint for next breath:
• If Vt_measured < Vt_target → increase flow
• If Vt_measured > Vt_target → decrease flow
Pressure-Limited Volume Control
When pressure reaches limit → flow reduced → volume sacrificed for safety

## Slide 9
Pressure Support Mechanism
Inspiration Support State
When patient trigger detected:
LLC Insufflation Pressure = Insp Pressure + Insp Support Pressure
Expiration Support State
When patient trigger detected:
LLC Insufflation Pressure = PEEP + Exp Support Pressure
Benefits:
Rewards patient inspiratory effort
Reduces work of breathing
Common in PSV (Pressure Support Ventilation) modes

## Slide 10
Inspiratory Pause Implementation
Purpose:
Plateau pressure measurement
Assessment of static lung compliance
Equilibration of airway and alveolar pressures
The Mechanism: Elegant Setpoint Manipulation
Bias Flow → 0
PEEP Setpoint → 100 cmH2O
Result: System is mechanically locked
No gas flows in (zero bias flow)
No gas escapes (exp valve closed, actual pressure << 100 cmH2O)
Volume held constant without special pause mode logic

## Slide 11
Design Advantages
Modularity
Clear separation of responsibilities across layers
No Direct Commands
Simplifies interfaces and reduces coupling
Extensibility
New ventilation modes via timing/parameter configuration
Patient Synchronization
Built into state machine architecture
Safety
Multiple protection layers (LLC pressure limiting, HLC limits)
Adaptability
Breath-to-breath learning handles patient variability

## Slide 12
Ventilator Control Architecture
Questions?