## Slide 1: LundaLoggern
En datalogger
Åke 2026-02-27

## Slide 2: LundaLoggern

## Slide 3: LundaLoggern

## Slide 5: Layer 1 – Characteristics
Each controller operates independently
Handles non-linearities in sensors and actuators
Provides fast, stable control loops (100–1000 Hz)
No knowledge of breathing phases or ventilation modes
Responsibilities
  Sensor scaling and compensation
  Actuator linearization
  Setpoint tracking with PID (or more advanced) control
Fast, independent control loops

## Slide 6: Layer 2
Low-Level Controller (LLC)
Breathing Phase Control

## Slide 7: LLC – Breathing Phase States
Two alternating states controlled by setpoints and triggers
Insufflation (Inspiration)
  Delivers pressure/flow to patient
  Controls gas via O₂ and Air flow controllers
  Maintains target inspiratory pressure
  Pressure limiting: reduces flow when pressure approaches limit
Exsufflation (Expiration)
  Releases pressure/flow from patient
  Controls expiratory valve for pressure release
  Maintains PEEP (baseline pressure)

## Slide 8: LLC – Parameters (Set by HLC)
FiO₂ – Fraction of inspired oxygen (0.21 – 1.0)
Insufflation Pressure – Target during insufflation
Pressure Limit – Maximum allowed airway pressure
Exsufflation Pressure (PEEP) – Target during exsufflation
Insufflation Flow – Target flow rate during insufflation
Bias Flow – Continuous flow during exsufflation
Max Inspiratory Flow – Maximum allowed inspiratory flow
Flow Trigger Threshold – Detects patient effort (flow)
Pressure Trigger Threshold – Detects patient effort (pressure)
All parameters written by HLC – no explicit commands

## Slide 9: LLC – Transitions & Communication
Purely reactive controller – no explicit commands from HLC
State Transitions
  Patient Effort Detection
    Monitors flow & pressure vs trigger thresholds
    Inspiratory effort → triggers insufflation
    Expiratory effort → triggers exsufflation
  Setpoint-Induced Transitions
    HLC modifies setpoints → trigger criteria fulfilled
    Example: HLC sets PEEP=100 cmH₂O → LLC shifts to exsufflation
Signals to HLC
  State change events (insuff ↔ exsuff)
  Measured tidal volume (integrated flow)
  Pressure limit events

## Slide 10: Layer 3
High-Level Controller (HLC)
Ventilation Mode Control

## Slide 11: HLC – 8-State Sequential Machine
Two phases × four states per phase
Inspiration Sequence
  1. Non-triggerable – Fixed inspiration, triggers ignored
  2. Support – Patient can trigger pressure support
  3. Synch – Patient-synchronized transition (default: 0)
  4. Pause – Inspiratory hold for plateau measurement
Expiration Sequence
  5. Non-triggerable – Fixed expiration, triggers ignored
  6. Support – Patient can trigger expiratory support (rare)
  7. Synch – Patient-synchronized transition (default: 0)
  8. Pause – Expiratory hold for measurements

## Slide 12: HLC – Inspiration States
1. Non-triggerable
  Delivers mandatory breath
  Patient triggers are ignored
  Duration set by timing parameters
2. Support
  Patient can trigger additional inspiratory support pressure
  Allows patient to augment mechanical breath
3. Synch (Synchronization)
  LLC exsufflation signal → HLC advances to Pause
  Allows patient to terminate inspiration
4. Pause (Inspiratory Plateau)
  Holds delivered volume – no gas added or released
  Used for plateau pressure measurement
States 1–4 of the ventilation cycle

## Slide 13: HLC – Expiration States
5. Non-triggerable
  Ensures minimum expiration time
  Patient triggers are ignored
6. Support
  Patient can trigger expiratory support pressure
  Rare in typical ventilation modes
7. Synch (Synchronization)
  LLC insufflation signal → HLC advances to Pause
  Allows patient to initiate next breath
8. Pause (Expiratory)
  Maintains expiratory state
  Used for breath holds or measurements
States 5–8 of the ventilation cycle

## Slide 14: HLC – State Transition Logic
States advance based on:
  Time elapsed – each state has a pre-calculated absolute duration
  Zero-duration skipping – states with time = 0 are immediately skipped
  Patient synchronization – during Synch states, LLC signals trigger advance
    Insp Synch: LLC exsufflation signal → advance to Insp Pause
    Exp Synch: LLC insufflation signal → advance to Exp Pause
Key design principle
  The HLC operates exclusively on absolute times (ms/s)
  It does NOT interpret RR, I:E ratio, pause %, or any user parameters
  All parameter-to-time conversion is done by the UI Layer above
Simple, deterministic transitions – no parameter interpretation

## Slide 15: HLC – Timing Interface
Eight absolute durations – the only timing inputs to the HLC
Inspiration phase
  Time_Insp_NonTrig – Mandatory inspiration duration
  Time_Insp_Support – Inspiration support window
  Time_Insp_Synch – Max wait for patient-synchronized termination
  Time_Insp_Pause – Inspiratory pause / plateau hold
Expiration phase
  Time_Exp_NonTrig – Mandatory expiration duration
  Time_Exp_Support – Expiration support window
  Time_Exp_Synch – Max wait for patient-synchronized termination
  Time_Exp_Pause – Expiratory pause / hold
Sum of all 8 durations = one complete breathing cycle
  Active Insp. Time = NonTrig + Support + Synch (pause excluded for Vt calc)

## Slide 16: User Interface Layer
Translates clinician parameters into HLC absolute times
UI Layer accepts user parameters
  Respiratory Rate (RR)
  I:E Ratio
  Inspiratory Pause %
  Expiratory Pause %
  Non-triggerable / Support / Synch fractions
  Tidal Volume, FiO₂, pressures, etc.
UI Layer solves the equations
  Cycle Time = 60 / RR
  Insp Time = Cycle × I/(I+E)
  Exp Time = Cycle × E/(I+E)
  Distributes durations across 8 states
  Writes absolute times + setpoints to HLC
Benefit: different UIs (touchscreen, remote,
automated protocols) share the same
deterministic HLC interface

## Slide 17: Tidal Volume Control
Breath-to-breath adaptive control with pressure safety
Initial Flow Calculation
  Insufflation Flow = Vt / Active Inspiration Time
  Active Insp. Time = Time(Non-trig) + Time(Support) + Time(Synch)
Breath-to-Breath Adaptation
  At each inspi→expi transition:
    Measure actual Vt_measured
    Calculate ΔVt = Vt_target − Vt_measured
    Gradually adjust insufflation flow setpoint
    Prevents oscillations, compensates for compliance changes
Pressure-Limited Volume Control
  LLC pressure controllers auto-reduce flow near Pressure Limit
  Vt_measured < Vt_target → volume sacrificed for safety
  Two modes: Volume-controlled / Pressure-limited volume-controlled

## Slide 18: Pressure Support Mechanism
Inspiration Support State
  When patient trigger detected during Insp Support state:
  LLC Insuff Pressure = Inspiratory Pressure + Insp Support Pressure
  Provides additional pressure boost rewarding inspiratory effort
  Common in PSV (Pressure Support Ventilation) modes
Expiration Support State
  When patient trigger detected during Exp Support state:
  LLC Insuff Pressure = PEEP + Exp Support Pressure
  Less common – available for specific clinical scenarios
Rewarding patient effort through pressure boosts

## Slide 19: Inspiratory Pause Implementation
Elegant setpoint manipulation – no special pause logic in LLC
Purpose
  Hold delivered volume for plateau pressure measurement
  Assess static lung compliance
  Equilibrate airway and alveolar pressures
Mechanism (setpoint manipulation)
  Bias Flow → 0 (no gas added by insufflation controllers)
  PEEP setpoint → 100 cmH₂O (forces LLC into exsufflation mode)
  Actual pressure ≪ 100 → expiratory valve stays closed
  Result: system is mechanically "locked"
    No gas flows in, no gas escapes
    Volume held constant

## Slide 20: HLC – Indirect Control Philosophy
HLC controls LLC by:
  Setting insufflation pressure based on state & triggers
  Setting pressure limit for safety
  Setting insufflation flow (Vt, timing, adaptation)
  Adjusting bias flow & PEEP for pause phases
  Adjusting trigger thresholds per HLC state
State-Dependent Setpoints (key states)
  Insp Non-trig: Insp Pressure | Normal PEEP | Normal Bias
  Insp Support (triggered): Insp Press + Support | Normal PEEP
  Insp Pause: N/A | PEEP=100 cmH₂O | Bias=0
  Exp states: N/A or PEEP+Support | Normal PEEP | Normal Bias
No explicit commands – only setpoint manipulation

## Slide 21: Communication Between Layers
Clean interfaces with minimal coupling
HLC → LLC (Setpoints)
  Pressure setpoints (insp, PEEP, limit)
  Flow setpoints (insufflation, bias, max insp)
  Trigger thresholds (flow & pressure)
LLC → HLC (Signals)
  State change events (insufflation ↔ exsufflation)
  Measured tidal volume (integrated flow)
  Pressure limit events
LLC → Layer 1 (Setpoints)
  Individual flow controller setpoints
  Expiratory valve pressure setpoint

## Slide 22: Example Ventilation Modes
Various clinical modes via timing & parameter configuration
Volume Control (VC)
  Insp Non-trig delivers full breath
  Support & Synch skipped (time=0)
  Time-triggered, volume-cycled
Assist Control (AC)
  Backup mandatory breath timing
  Exp Synch active → patient triggering
  Full support on every breath
Pressure Support (PSV)
  Insp Support active with pressure
  Insp & Exp Synch active
  Fully patient-triggered & cycled
SIMV
  Mandatory (Non-trig) + Supported breaths
  Synch states coordinate with patient
  Combines mandatory & spontaneous

## Slide 23: Safety & Design Advantages
Safety Considerations
  Pressure limiting always active (barotrauma prevention)
  Non-trig states ensure minimum ventilation
  Gradual adaptation prevents sudden volume changes
  Setpoint sanity checks in HLC
  Watchdog timers per state
Design Advantages
  Modularity – clear layer responsibilities
  No direct commands – reduced coupling
  Extensibility – new modes via config only
  Patient synchronization built-in
  Breath-to-breath adaptability

## Slide 24: Implementation Notes
Execution Rates
  Layer 1 PID loops: 100 – 1000 Hz
  LLC: event-driven architecture
  HLC: fixed time base (10 – 50 Hz)
Parameter Updates
  Validate all parameter changes before applying
  Rate-limit parameter changes to prevent instability
  Log all setpoint changes for debugging & clinical review
Testing Strategy
  Test each layer independently first
  Verify LLC response to setpoint changes
  Verify HLC state sequences for all modes
  Patient simulator testing with varying compliance/resistance
Practical considerations for system design