## Slide 1: Three-Layer Ventilator Control Architecture
A modular, setpoint-driven approach
to medical ventilator control
v1.1 · 2026-02-01

## Slide 2: Architecture Overview
Three-layer hierarchy
  Layer 1 — Actuator Layer: fast PID loops for individual valves and blower
  Layer 2 — LLC: breathing phase state machine (insufflation / exsufflation)
  Layer 3 — HLC: ventilation mode logic, 8-state cycle
Core philosophy: indirect control through setpoint manipulation
  Higher layers never issue explicit commands — they change setpoints only
  Lower layers react automatically when trigger conditions are met
Update rates
  Layer 1: 100–1000 Hz Layer 2: ~100 Hz Layer 3: 10–50 Hz

## Slide 3: Layer 1: Actuator Controllers
Four independent PID controllers
  O2 Flow Controller — controls O2 valve
  Air Flow Controller — controls air valve, handles linearisation tables
  Expiratory Valve Controller — controls exp pressure / valve position
  Blower Flow Controller — controls blower flow rate
Key characteristics
  Each controller operates independently at the highest loop rate
  Handles sensor & actuator non-linearities via piecewise lookup tables
  No knowledge of breathing phases or ventilation modes
  All setpoints received from LLC — outputs go directly to hardware drivers

## Slide 4: Interface: LLC → Layer 1
LLC writes setpoints; Layer 1 controllers read them each cycle
During Insufflation
  Air flow setpoint [slm] — to Air Flow Controller
  O2 flow setpoint [slm] — to O2 Flow Controller (FiO2 determines split)
  Exp valve pressure [mbar] — held at high value to keep exp valve closed
During Exsufflation
  Air & O2 flow setpoints → 0 (or Bias Flow) — valves nearly closed
  Exp valve pressure setpoint = PEEP — exp valve opens to target pressure
Pressure limiting (active at all times)
  Layer 1 flow controllers reduce output when Paw approaches Pressure Limit
  LLC does not need to poll for this — Layer 1 handles it autonomously

## Slide 5: Layer 2: LLC — Breathing Phase Control
Insufflation (Inspiration)
  Delivers pressure / flow to patient via O2 and Air controllers
  Maintains target inspiratory pressure
  Pressure limiting: reduces flow when Paw approaches pressure limit
Exsufflation (Expiration)
  Controls expiratory valve to manage pressure release
  Maintains PEEP (positive end-expiratory pressure)
State transitions triggered by
  Patient effort — flow or pressure crosses trigger threshold
  Setpoint changes from HLC — may instantly fulfil trigger criteria

## Slide 6: LLC: Parameters Set by HLC
Pressure: Insufflation Pressure, Pressure Limit, PEEP
Flow: Insufflation Flow, Bias Flow, Max Inspiratory Flow
Oxygen: FiO2 (0.21 – 1.0)
Triggers: Flow Trigger Threshold, Pressure Trigger Threshold

## Slide 7: Interface: HLC ↔ LLC
HLC → LLC (setpoints, updated every HLC cycle)
  Insufflation Pressure [cmH2O / mbar]
  Pressure Limit [cmH2O / mbar] — hard safety ceiling
  PEEP [cmH2O / mbar] — exsufflation target pressure
  Insufflation Flow [slm] — adapted breath-to-breath for Vt
  Bias Flow [slm] — set to 0 during Insp Pause
  FiO2 [0.21 – 1.0]
  Flow Trigger Threshold [slm] — HLC adjusts per state
  Pressure Trigger Threshold [cmH2O] — HLC adjusts per state
LLC → HLC (events, asynchronous)
  State change event — fired on insufflation ↔ exsufflation transition
  Tidal volume report — integrated flow delivered during inspiration [ml]
  Pressure limit event — fired when Layer 1 pressure limiting was active

## Slide 8: Layer 3: HLC — 8-State Breathing Cycle
Inspiration sequence
  1. Non-triggerable — mandatory breath, patient triggers ignored
  2. Support — patient trigger adds Insp Support Pressure
  3. Insp Synch — LLC exsufflation event advances HLC (default: 0 s)
  4. Insp Pause — holds volume; Bias Flow = 0, PEEP = 100 cmH2O
Expiration sequence
  5. Non-triggerable — minimum expiration time, triggers ignored
  6. Support — patient trigger adds Exp Support Pressure
  7. Exp Synch — LLC insufflation event advances HLC (default: 0 s)
  8. Exp Pause — holds expiratory state for measurement

## Slide 9: HLC: State Transitions & Timing
States advance on
  Time elapsed — each state has a configured duration
  Zero-duration skip — states with time = 0 are immediately bypassed
  Patient sync — LLC state-change events trigger HLC advance in Synch states
Timing parameters
  Respiratory Rate (RR) Cycle time = 60 / RR seconds
  I:E Ratio Splits cycle into inspiration / expiration time
  Insp Pause Fraction Pause time = fraction x (60 / RR)
  Active Inspiration Time = Non-trig + Support + Synch (pause excluded)
Active Inspiration Time drives the tidal volume flow calculation

## Slide 10: Tidal Volume Control
Initial flow calculation
  Insufflation Flow = Vt_target / Active Inspiration Time
  Pause time excluded — no gas flows during plateau hold
Breath-to-breath adaptive control (at each insp → exp transition)
  Measure Vt_delivered (integrated flow over inspiration from LLC event)
  Error: dVt = Vt_target - Vt_delivered
  Gradually adjust Insufflation Flow setpoint for next breath
  Compensates for changing patient compliance and resistance
Pressure-limited mode
  LLC pressure limiting reduces flow — LLC reports pressure limit event to HLC
  HLC attempts to increase flow next breath; if persistent: volume sacrificed
  Patient safety (pressure) prioritised over volume target

## Slide 11: Pressure Support & Inspiratory Pause
Inspiration Support (on patient trigger during Insp Support state)
  HLC sets: Insufflation Pressure = Inspiratory Pressure + Insp Support Pressure
  Rewards patient effort; foundation of PSV mode
Expiration Support (on patient trigger during Exp Support state)
  HLC sets: Insufflation Pressure = PEEP + Exp Support Pressure
Inspiratory Pause — lock using existing setpoints only
  HLC sets: Bias Flow = 0 and PEEP = 100 cmH2O
  PEEP 100 cmH2O forces LLC into exsufflation mode
  Actual Paw << 100 cmH2O => expiratory valve stays physically closed
  Result: zero inflow + valve closed = volume held constant
  No special pause logic needed in LLC

## Slide 12: Indirect Control: State-Dependent Setpoints
HLC controls LLC exclusively by changing setpoints — never by explicit command
HLC State Insp Pressure PEEP Bias Flow
  Insp Non-trig Insp Pressure Normal PEEP Normal
  Insp Support Insp P + Support * Normal PEEP Normal
  Insp Synch Insp Pressure Normal PEEP Normal
  Insp Pause — 100 cmH2O (lock) 0 (lock)
  Exp Non-trig — PEEP Normal
  Exp Support PEEP + Support * PEEP Normal
  Exp Pause — PEEP Normal
* Only applied when patient trigger is detected in that state
Setpoint change → LLC trigger criteria may be instantly fulfilled
  LLC transitions automatically — no command needed from HLC

## Slide 13: Clinical Ventilation Modes
Volume Control (VC)
  Insp Non-trig: full mandatory breath
  Support / Synch: skipped (0 s)
  Result: time-triggered, volume-cycled
Assist Control (AC)
  Insp Non-trig: backup mandatory timing
  Exp Synch: patient can trigger next breath
  Full support on every breath
Pressure Support (PSV)
  Insp & Exp Synch states active
  Insp Support: pressure boost on trigger
  Fully patient-triggered and patient-cycled
SIMV
  Mandatory breaths via Non-trig states
  Supported breaths via Support states
  Synch states coordinate with patient effort

## Slide 14: Safety & Design Advantages
Safety layers
  Pressure limiting always active in Layer 1 — prevents barotrauma
  Non-trig states guarantee minimum ventilation if patient effort fails
  Gradual Vt adaptation prevents sudden volume changes
  Setpoint sanity checks in HLC before values reach LLC
  Watchdog timers on each HLC state (maximum duration limits)
Design advantages
  Modularity — clear single responsibility per layer
  No direct commands — simpler interface, lower coupling between layers
  Extensibility — new ventilation modes by timing & parameter config only
  Patient synchronisation built into the state machine
  Breath-to-breath learning handles varying compliance and resistance